## Slide 1: Calibration
Ron Maddalena
NRAO – Green Bank
November 2012

## Slide 2
Receiver calibration sources allow us to convert the backend’s detected voltages to the intensity the signal had at the point in the system where the calibration signal is injected.

## Slide 3: Reference observations
Difference a signal observation with a reference observation
Types of reference observations
  Frequency Switching
    In or Out-of-band
  Position Switching
  Beam Switching
    Move Subreflector
    Receiver beam-switch
  Dual-Beam Nodding
    Move telescope
    Move Subreflector

## Slide 4: Position-Switched (On-Off) Observing

## Slide 5: Typical Position-Switched Calibration Equation for a Point Source
A(Elev,t) = Air Mass
τ(ν,t) = Atmospheric Zenith Opacity
Tcal = Noise Diode Temperature
Area = Physical area of the telescope
ηA (ν,Elev) =Aperture efficiency (point sources)
TA (ν) = Source Antenna Temperature
S(ν) = Source Flux Density
Sig(ν), Ref (ν) = Data taken on source and
on blank sky (in units backend counts)
On,Off = Data taken with the noise diode on and off
Tsys = System Temperature averaged over bandwidth

## Slide 6: Position-Switched Calibration Equation
Sources of uncertainties
10-15% accuracy have been the ‘standard’
Usually, errors in Tcal dominate
Goal: To achieve 5% calibration accuracy without a significant observing overhead.

## Slide 7: Air Mass Estimates
Depends upon density and index of refraction as a function of height
But, how can one get this information?

## Slide 8: Air Mass Estimate
Air Mass traditionally modeled as 1/sin(Elev)
For 1% calibration accuracy, must use a better model below 15 deg.
Good to 1 deg
Use 1/sin(Elev) above 60 deg
Coefficients are site specific, at some low level

## Slide 9: Typical Position-Switched Calibration Equation
A(Elev,t) = Air Mass
τ(ν,t) = Atmospheric Zenith Opacity
Tcal = Noise Diode Temperature
Area = Physical area of the telescope
ηA (ν,Elev) =Aperture efficiency (point sources)
TA (ν) = Source Antenna Temperature
S(ν) = Source Flux Density
Sig(ν), Ref (ν) = Data taken on source and
on blank sky (in units backend counts)
On,Off = Data taken with the noise diode on and off
Tsys = System Temperature averaged over bandwidth

## Slide 10: Opacities from the various components
Dry Air Continuum
Water Continuum
Oxygen Line
Water Line
Hydrosols

## Slide 11: Opacities from the various components
Total Opacity
gfs3_c27_1190268000.buf

## Slide 12: Determining Opacities
Slope ~ TATM

## Slide 13: Typical Position-Switched Calibration Equation
A(Elev,t) = Air Mass
τ(ν,t) = Atmospheric Zenith Opacity
Tcal = Noise Diode Temperature
Area = Physical area of the telescope
ηA (ν,Elev) =Aperture efficiency (point sources)
TA (ν) = Source Antenna Temperature
S(ν) = Source Flux Density
Sig(ν), Ref (ν) = Data taken on source and
on blank sky (in units backend counts)
On,Off = Data taken with the noise diode on and off
Tsys = System Temperature averaged over bandwidth

## Slide 14: Determining TCal from hot-cold load measurements in the lab
Place black bodies (absorbers) of two known temperatures in front of the feed and record detected voltages.
    VHot_Off = g * THot
    VCold_Off = g * TCold
    VCold_On = g * (TCold + TCal)
g and TCal are unknown

## Slide 15: Determining TCal from hot-cold load measurements in the lab
Course frequency resolution
Uncertainties in load temperatures
Are the absorbers black bodies?
Detector linearities (300 & 75 K)
Lab TCal may be off by 10%
So… all good observers perform their own astronomical calibration observation

## Slide 16: Noise Diode Estimates
Instead, we recommend an On-Off observation
  Use a point source with known flux -- polarization should be low or understood
  Use the same exact hardware, exact setup as your observation. (i.e., don’t use your continuum pointing data to calibrate your line observations.)
  Observations take ~5 minutes per observing run
  Staff take about 2 hrs to measure the complete band of a high-frequency, multi-beam receiver.
  Resolution sufficient: 1 MHz, sometimes better
  Accuracy of ~ 1%, mostly systematics.

## Slide 17: Noise Diode Estimates
Remove Averaging, Solve for Tcal

## Slide 18: Noise Diode Estimates

## Slide 19: Typical Position-Switched Calibration Equation
A(Elev,t) = Air Mass
τ(ν,t) = Atmospheric Zenith Opacity
Tcal = Noise Diode Temperature
Area = Physical area of the telescope
ηA (ν,Elev) =Aperture efficiency (point sources)
TA (ν) = Source Antenna Temperature
S(ν) = Source Flux Density
Sig(ν), Ref (ν) = Data taken on source and
on blank sky (in units backend counts)
On,Off = Data taken with the noise diode on and off
Tsys = System Temperature averaged over bandwidth

## Slide 20: Telescope efficiencies – Part 1

## Slide 21: GBT Gain Curve
Ruze Equation – Surface errors

## Slide 22: GBT Gain Curve

## Slide 23: Non-linearity
If system is linear,
  Pout = B * Pin
  (SigOn-SigOff) – (RefOn-RefOff ) = 0
Model the response curve to 2nd order:
  Pout = B * Pin + C * Pin2
Our ‘On-Off’ observations of a calibrator provide:
  Four measured quantities: Refoff, Refon, Sigoff, Sigon
  TA From catalog
  Four desired quantities: B, C, Tcal, Tsys
It’s easy to show that:
  C = [(Sigon- Sigoff )-(Refon- Refoff)]/(2TATcal)
Thus:
  Can determine if system is sufficiently linear
  Can correct to 2nd order if it is not

## Slide 24: Non-linearity
(SigOn-SigOff) – (RefOn-RefOff )

## Slide 25: Reference observations
Difference a signal observation with a reference observation
Types of reference observations
  Frequency Switching
    In or Out-of-band
  Position Switching
  Beam Switching
    Move Subreflector
    Receiver beam-switch
  Dual-Beam Nodding
    Move telescope
    Move Subreflector

## Slide 26: Position switching
Move the telescope between a signal and reference position
  Overhead
  ½ time spent off source
Difference the two spectra
Assumes shape of gain/bandpass doesn’t change between the two observations.
For strong sources, must contend with dynamic range and linearity restrictions.

## Slide 27: Frequency switching
Eliminates bandpass shape from components after the mixer
Leaves the derivative of the bandpass shape from components before the mixer.

## Slide 28: In-Band Frequency Switching

## Slide 29: Out-Of-Band Frequency Switching

## Slide 30: Beam switching – Internal switch
Difference spectra eliminates any contributions to the bandpass from after the switch
Residual will be the difference in bandpass shapes from all hardware in front of the switch.
Low overhead but ½ time spent off source

## Slide 31: Atmosphere is in the near field
Common to all feeds in a multi-feed receiver

## Slide 32: Beam Switching – Subreflector or tertiary mirror
Optical aberrations
Difference in spillover/ground pickup
Removes any ‘fast’ gain/bandpass changes
Low overhead. ½ time spent off source

## Slide 33: Nodding with dual-beam receivers - Telescope motion
Optical aberrations
Difference in spillover/ground pickup
Removes any ‘fast’ gain/bandpass changes
Overhead from moving the telescope. All the time is spent on source

## Slide 34: Nodding with dual-beam receivers - Subreflector or tertiary mirror
Optical aberrations
Difference in spillover/ground pickup
Removes any ‘fast’ gain/bandpass changes
Low overhead. All the time is spent on source

## Slide 35: References
Rohlfs & Wilson, “Tools of Radio Astronomy”
Stanimrovis et al, “Single-Dish Radio Astronomy: Techniques and Practices”
Baars, 1973, IEEE Trans Antennas Propagat, Vol AP-21, no. 4, pp 461-474
Kutner & Ulich, 1981, Astronomica Journal, Vol 250,pp 341-348
Winkel, Kraus, & Bach, 2012, Astronomy & Astrophysics, vol. 540.